INSTITUTE OF MICROBIAL TECHNOLOGY (IMTECH),

(Council of Scientific & Industrial Research)

Sector 39-A, Chandigarh -160036 (India)

## Ph: 0091-172-2690056/2695225/6/7 ext. 331/334, fax:2 690585/2690632,

## e-mail: purchase@imtech.res.in

Top of Form

| No: PUR/2009-10/IND11122/KK | Date: 19/01/2010 |
| --- | --- |

Sub : Request for Pro-forma Invoice Dear Sir This Research Institute is interested in purchasing of the items mentioned below. You are therefore requested to forward a pro-forma invoice along with your terms & conditions of supply & payment. Please provide a copy of the technical literature and the specifications of these items in a sealed envelope. Please read carefully the terms and conditions mentioned overleaf and submit your pro-forma invoice/quotation accordingly.

| SNo. | Name of Item | Catalog No | Qty |
| --- | --- | --- | --- |
| 1 | High speed refrigerated centrifuge, specifications attached |  | 1 |

Important Please mention our reference number on the envelop and address all correspondence to Director IMT Chandigarh only. Please quote strictly as per our specifications. Please attach compliance sheet Proforma Invoice received after due date will not be considered under any circumstances. Please note that your Quotation should reach us latest by 02/02/2010 opened on 03/02/2010 Your Quotation must be duly stamped and reach directly to IMT from you only. If quotation is submitted by Indian representative/agent then they must have to produce a authority certificate of principal party for quotating the price Other wise it will be very difficult at our end to consider your quotation. EMD Rs.50,000/- must be deposited along with Technical Bid Quotation must be submitted under two bid system (1) Tech. Bid (2) Price Bid

Stores & Purchase Officer.

Bottom of Form

# (1.)	PUR/2009-10/IND11122

High Speed Refrigerated Centrifuge

Specifications :

- Speed range : 500 – 21,000 rpm

- Speed control accuracy : +/- 50 rpm

- Maximum RCF : 50,000 x g

- Capacity :  4 Liters

- Temperature set range : -10 to +40 °C (with 1 °C increments)

- Temperature control : +2 to +40 °C

- Temperature accuracy : +/-2 °C

- Ambient temperature range : 15 – 35 °C

- CFC free refrigeration system

- Microprocessor controlled,  Delayed start and stop function

- Run time (hr/min) 99/59; Hold

- Voltage : 220-240V/50-60 Hz.

- A comprehensive 1 year warranty and, and one year additional warranty.

- At least 3 year warranty on the refrigeration system.

Optionals  to  be quoted separately :

Rotors :

- 6 x 500 ml

- 6 x 250 ml

- 14 x 50 ml

- 24x 16 ml

- 4 x 1,000 ml

- Other suitable rotors for volumes ranging from 1-50 ml and adapters to be used  may be quoted separately.

- Items needed for maintenance of the centrifuge

General conditions:

- The vendor should provide point-wise compliance statement in tabular form as given in tender document, clearly mentioning the compliance/deviation.

- All relevant literature in support of compliance statement should be provided.

- Details of AMC charges after warranty period

- List of users in other organizations

- Service manual

TERMS & CONDITIONS

Commercial

1) Cost                  : 	Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment.

2) Insurance         :	Consignment should be insured through National Insurance               Company of India or is associate in your country.

3) Freight              :	The dispatch of the material should be through Air India preferably.

4) Agency             : 	Mention the   percentage   of  Indian  Agency   Commission  payable  to

Commission                    commission Indian Agent in INR and also mention the name of your   authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end.

5) Validity             : 	The validity of your pro forma invoice should be less than 180 days from the date of opening.

6) Warranty           : 	Mention the warranty period of equipment after satisfactory   installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above

General Condition

- Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market.

- Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization.

- Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years.

- Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required.

- Provide the list of the users in India of the same equipment along with their contact / email address.

- Enclose a copy of agreement with your authorized Indian Agent.

FAX : 0091-172-2690056, 2690632			                                   email : purchase@imtech.res.in

TEL  : 0091-0172-2690056, 2636680 - 94     	                                            website : www.imtech.res.in

INSTITUTE OF MICROBIAL TECHNOLOGY (IMTECH),

(Council of Scientific & Industrial Research)

Sector 39-A, Chandigarh -160036 (India)

## Ph: 0091-172-2690056/2695225/6/7 ext. 331/334, fax:2 690585/2690632,

## e-mail: purchase@imtech.res.in

High Speed Refrigerated Centrifuge

Specifications :

- Speed range : 500 – 21,000 rpm

- Speed control accuracy : +/- 50 rpm

- Maximum RCF : 50,000 x g

- Capacity :  4 Liters

- Temperature set range : -10 to +40 °C (with 1 °C increments)

- Temperature control : +2 to +40 °C

- Temperature accuracy : +/-2 °C

- Ambient temperature range : 15 – 35 °C

- CFC free refrigeration system

- Microprocessor controlled,  Delayed start and stop function

- Run time (hr/min) 99/59; Hold

- Voltage : 220-240V/50-60 Hz.

- A comprehensive 1 year warranty and, and one year additional warranty.

- At least 3 year warranty on the refrigeration system.

Optionals  to  be quoted separately :

Rotors :

- 6 x 500 ml

- 6 x 250 ml

- 14 x 50 ml

- 24x 16 ml

- 4 x 1,000 ml

- Other suitable rotors for volumes ranging from 1-50 ml and adapters to be used  may be quoted separately.

- Items needed for maintenance of the centrifuge

General conditions:

- The vendor should provide point-wise compliance statement in tabular form as given in tender document, clearly mentioning the compliance/deviation.

- All relevant literature in support of compliance statement should be provided.

- Details of AMC charges after warranty period

- List of users in other organizations

- Service manual

Store & Purchase Officer

(2)   PUR/2009-10/IND10236

Laboratory Scale Fermenter

Specifications

- Pur/2009-10/IND10237

Fermenter, Capacity 18-20 L/40-50 L

Specifications:

TERMS & CONDITIONS

Commercial

1) Cost                  : 	Mention the price of the material in the FOB/FCA including inland freight value only. Please note that being a Govt. Organization we cannot make Advance Payment.

2) Insurance         :	Consignment should be insured through National Insurance               Company of India or is associate in your country.

3) Freight              :	The dispatch of the material should be through Air India preferably.

4) Agency             : 	Mention the   percentage   of  Indian  Agency   Commission  payable  to

Commission                    commission Indian Agent in INR and also mention the name of your   authorized Indian Agent/Representative, Pro-forma invoice without mentioning Indian Agency Commission will not be considered at our end.

5) Validity             : 	The validity of your pro forma invoice should be less than 180 days from the date of opening.

6) Warranty           : 	Mention the warranty period of equipment after satisfactory   installation. 10% Bank Guarantee is to be submitted by you/your Agent for the Equipments costing Rs.25 Lac & above

General Condition

- Zerox copies of latest price list applicable list may be enclosed applicable in Indian Market.

- Ensure that quoted price will not be higher to any other customers in India to whom this particular item have been sold particularly to govt. of India Organization.

- Sent a certificate that the equipment is of latest technology & will not be obsolete within 5 years.

- Ensure that all essential spares parts of this equipment will available in India with your Indian Agent as and when required.

- Provide the list of the users in India of the same equipment along with their contact / email address.

- Enclose a copy of agreement with your authorized Indian Agent.

FAX : 0091-172-2690056, 2690632			                                   email : purchase@imtech.res.in

TEL  : 0091-0172-2690056, 2636680 - 94     	                                            website : www.imtech.res.in